## Slide 1: UK feedback on MQO
Presented by: John Stedman, Daniel Brookes, Brian Stacey, Keith Vincent, Emily Connolly
10 April 2013

## Slide 2: Outline
Current view on the proposed MQOs
Other aspects covered in accompanying presentation
  MQO formulation
  NO2 measurement uncertainty
  Fitting procedure
  Application to NO2
  PM measurement uncertainty
  Fitting procedure
  Application to PM10
  Conclusions and recommendations
Views of the UK Competent Authorities

## Slide 3: Current view of the proposed MQOs
  NO2:
    Uncertainty budget for hourly measurements largely reasonable.
    Less happy with the application to annual means in terms of the cancelling of random errors, specifically the lack of fit/linearity component.
    Overall we think that the latest version of the coefficients for annual mean NO2 are still a bit too stringent at low concentrations.

## Slide 4: Current views of the proposed MQOs
  PM10:
    There are a set of coefficients defined for each measurement type in early versions of the paper which do not appear in the latest paper, although data for all measurements appears in Figures.
    In the Delta tool and in the latest paper the most stringent coefficients (gravimetric measurement based) have been carried through.
    The resulting model DQOs (gravimetric measurement based) are too stringent at all concentrations for annual mean PM10 on this basis.
    TEOM (presumably FDMS) coefficients from an earlier version of the paper result in more generous uncertainty limits.

## Slide 5: MQO formulation
T2012 proposed MQO:
T2013 Part I: Simplified formulation for RMSU
Proportional component
Non-Proportional component

## Slide 6: MQO formulation
T2013 Part II: MQO for annual average results
T2013 Part II: Extension of uncertainty formulation for time averaging
  Introduction of Np and Nnp to account for autocorrelation
  Dropping of σ2

## Slide 7: MQO formulation
Shouldn’t this be?
T2013 Part II: Drops σ2 using the substitution:
Only valid if Np* is const. and independent of xm or a constant function of xm such that Np* = f(xm) = const.

## Slide 8: MQO formulation
T2013 Part II: However... using NO2 monitoring data from 80 UK national network monitoring sites for the year 2010

## Slide 9: NO2 measurement uncertainty
Based on GUM methodology, type B uncertainty
  Broadly happy but...
  Cancelling of random errors, specifically the lack of fit/linearity component is unreasonable
994 urban stations in AirBase, 2009 data
  Representative of all years?

## Slide 10: NO2 measurement uncertainty
  T2013 Part II: Table B.1
  Lack of fit, linearity component is the largest component of NO2 uncertainty budget
  Is this uncertainty component normally distributed and 100% random?
  Not the case:

## Slide 11: Fitting procedure
Linear fit of uc(xi)2 vs xi2 for hourly, uc(xm)2 vs xm2 for annual (so missing σ2 – should be uc(xm)2 vs xm2 + σ2)
Constant coefficient RV a reference value set at hourly LV
Constant coefficients urRV and α calculated from linear fit of hourly NO2 data
Constant coefficients Np* and Nnp calculated from linear fit of annual average NO2 data, holding urRV, α and RV constant
2 μgm-3 offset applied to annual fit to avoid underestimation of uncertainty at low concentrations

## Slide 12: Fitting procedure:
Residuals for hourly NO2 fit show non-linearity, overestimate uncertainty
But estimating maximum uncertainty so overestimate ok?
Annual NO2 fit also shows non-linearity
Tendency to underestimate:
Hence 2 μgm-3 offset applied, and Np and Nnp re-calculated to estimate maximum uncertainty
Hourly values
Yearly values

## Slide 13: Fitting procedure:
Explanation for non linearity at lower concentrations suggested as resulting from < 1 correlation between NO and NOx at low NO2 (Gerboles et al, 2003)
Sensitivity of the fit coefficients to the determination of the gradient and intercept
Sensitivity to the underlying measurement data so will be sensitive to year to year variations in observed concentrations
Approximation of measurement uncertainty, attempting to define maximum uncertainty

## Slide 14: Application to NO2: PCM model results for 2010
  Parameter values in V3.0 (left) are not consistent with the paper circulated on 5 March 2013 (right)

## Slide 15: Application to NO2: PCM model 2010 and 2011
  Model performance varies from year to year
    Using parameters from 5 March 2013 paper

## Slide 16: Application to NO2: Sensitivity to inclusion of σ2
  With (left) and without σ2 term (right)
    Using parameters from 5 March 2013 paper

## Slide 17: PM10 measurement uncertainty
  GUM methodology, type B uncertainty:
  T2013 Part II: Table C.1
  Should be referencing the new prEN12341 standard
  u flow calibration – 1.7% in the new EN12341
  u mba balance calibration – 0.24ug/m3 in the new EN12341 (25/(3)0.5 = 0.24)
However, GUM method not applied:
  T2013 Part II: Appendix C – Limitations to estimate PM measurement uncertainty

## Slide 18: PM10 measurement uncertainty
Instead an approach based on GDE (2010) method for PM10 measurement uncertainty estimation
  Calibration chain: Demonstration of equivalence with gravimetric standard => transfer standard => Demonstration of equivalence with transfer standard
  Measurement uncertainty increases along calibration chain
  GDE method means measurement uncertainty defined under limited conditions => representative across Europe?

## Slide 19: PM10 measurement uncertainty
  Historically little evidence for demonstration of ongoing equivalence.
  Efforts underway to improve quantification of PM measurement uncertainty:
  WG15 working on quantification of uncertainty associated with filter media
  Evidence to feed into a new measurement standard

## Slide 20: PM10 measurement uncertainty
    Comparisons show large variation in the relationship between measurement types

## Slide 21: PM10 measurement uncertainty
  Previous versions of the paper: coefficients presented for gravimetric, teom and beta ray methods
  Current paper only presents coefficients for gravimetric which tend to be much more stringent.
  Uncertainty criteria applied should be appropriate to the measurement being compared:
  Most of the UK network is TEOM.

## Slide 22: Delta V3.0: PCM PM10 in 2010
  Using parameter values from 5 March paper (left)
  Using parameters for TEOM (FDMS) (right)

## Slide 23: Conclusions and recommendations
Model DQO Journal papers:
  What is the process to go from journal papers to technical guidance for MS?
Revising the requirements for reporting that are presently within the AQD?
  Is it proposed that this new method completely replaces the existing text in Annex I?
  Our previous understanding was to include a reference to Commission Guidance on model DQO in a revised AQD legal text and that this would then be developed by FAIRMODE.
  We now do not expect proposals for a new AQD for several years.
  How should this be taken forwards?
  How can we comply with the existing text in the interim once a new method is established but before the AQD is changed?

## Slide 24: Conclusions and recommendations
Other complications:
  Spatial representativity. ‘The measurements that have to be selected for comparison with modelling results shall be representative of the scale covered by the model.’
  Developments in quantification of measurement uncertainty
  Any revisions to the fit will lead to new coefficients to apply, new versions of Delta tool

## Slide 25: Concluding remarks
Need to decide whether these formulations are fit for use
The Directive defines model and measurement in the vicinity of the limit value
Implication of a burden in formulating measurement uncertainty at values other than the limit value if we adopt this approach